# Regulatory, Clinical Safety & Data Privacy Compliance Audit

Target System: LifeMetrics Ambient Intelligence System

Authors:

Amani Wanjiru, Security & Compliance Reviewer

Renata Cruz, Data Privacy Engineer

Target Audience: Clinical Safety Officers, DPOs, Executive Leadership

Reviewers: Sade (Product Strategy), Arjun (Solution Architecture)

Status: Compliance Specification v1.0

---

## 1. Executive Summary & Audit Scope

This audit outlines the regulatory, clinical safety, and data privacy framework required to deploy the LifeMetric Ambient Intelligence System in independent elderly care facilities and NHS clinical virtual wards.

By utilizing passive, camera-free mmWave radar, LWIR thermal grids, and acoustic micro-signals, the system solves the "Privacy Paradox" at the hardware level. However, because it extracts vital signs (Heart Rate, Respiratory Rate) and predicts physical events (Falls), it triggers high-stakes clinical and privacy compliance mandates under UK, EU, and US jurisdictions.

---

## 2. Clinical Safety Framework: UK DCB0129 Compliance

Author: Amani Wanjiru

In the United Kingdom, any health-tech software deployed within health or social care environments must comply with DCB0129 (Clinical Safety Risk Management System).

### 2.1 Clinical Safety Management System (CSMS)

To meet NHS mandates, LifeMetric must appoint a registered clinician as a Clinical Safety Officer (CSO) to establish and oversee the CSMS.

1.  Risk Analysis & Risk Hazard Log: The CSO must maintain a live, rolling Clinical Risk Hazard Log.

2.  Clinical Safety Case Report (CSCR): Prior to deployment, a CSCR must be compiled and submitted to NHS commissioners, detailing the mitigations of all identified clinical safety hazards.

### 2.2 Key Risk Hazard Log & Engineering Mitigations

| Hazard ID | Clinical Hazard | Severity | Mitigation Strategy |
| --- | --- | --- | --- |
| CL-HZ.1 | False Negative (Unreported Fall): System fails to detect an actual fall, leading to delayed medical intervention ("The Long-Lie"). | Critical | Implement Multi-Modal Sensor-Fusion. mmWave posture trajectory validation is cross-referenced in-memory with secondary LWIR thermal centroids and MEMS acoustic thud frequencies ($20\text{Hz}$–$200\text{Hz}$), maintaining sensitivity $>90\%$. |
| CL-HZ.2 | False Positive (Alert Fatigue): High volume of false fall alerts (caused by falling water, steam, or pet movements) causes nursing staff to disable alerts. | High | Apply background environmental clutter subtraction during local calibration. Use LWIR thermal mapping to isolate and reject biological heat centroids representing pets. |
| CL-HZ.3 | Latency-Driven Fall Alert Lag: Alert arrives at caregivers' badges $>5$ seconds post-fall, rendering predictive interventions useless. | High | Enforce a strict $<200\text{ms}$ low-latency edge-to-cloud signaling pipeline, utilizing binary UDP/WebSocket payloads and AWS RDS Proxy connection pooling. |

---

## 3. US FDA Class II Medical Device Certification Pathway

Author: Amani Wanjiru

Because LifeMetric passively monitors heart rate and respiratory rate, the device is classified as Software as a Medical Device (SaMD) under US FDA regulations.

### 3.1 Regulatory Classification Mapping

Classification Name: Cardiac/Respiration Rate Monitor (Passive).

Regulation Number: 21 CFR 870.2300 (or equivalent under SaMD guidelines).

Product Code: MJO / DRG.

Device Class: Class II (Special Controls).

Pathway: Pre-Market Notification [510(k)] showing substantial equivalence to an existing predicate device, or De Novo request if substantial equivalence cannot be proved.

### 3.2 Key FDA Compliance Requirements

1.  ISO 13485 QMS Alignment: The manufacturing and software development processes must operate under a Quality Management System (QMS) aligned with ISO 13485 standards.

2.  Clinical Validation Trials: Conduct clinical trials with a cohort of at least 150 participants to validate the accuracy of mmWave heart and respiratory rate extraction against gold-standard contact ECG and spirometry devices.

3.  Human Factors & Usability (ANSI/AAMI HE75): Document exhaustive usability testing to prove that clinical staff can interpret the "Actionable Wellbeing Summaries" dashboard without risk of clinical misinterpretation.

---

## 4. Global Data Privacy Compliance

Author: Renata Cruz

### 4.1 GDPR / UK GDPR Compliance Boundary

LifeMetric processes special category health data under the General Data Protection Regulation (GDPR) Article 9.

1.  Data Minimization (GDPR Article 5): The system implements data minimization by design. Because mmWave radar and LWIR sensors generate unstructured point-clouds and low-res grid arrays rather than identifiable images or video, no biometric PII is ever recorded, stored, or transmitted.

2.  Lawful Basis for Processing (Articles 6 & 9):

Lawful Basis (Article 6): Performance of a contract or legitimate interests (care delivery).

Special Category Basis (Article 9): Explicit, written, and unambiguous consent from the resident or their legally appointed guardian/power of attorney.

3.  The "Right to be Forgotten" (Article 17): Although the physical device maintains no storage partition, the cloud database stores historical vital trend logs. The system must support immediate, permanent purging of a patient's historical trends upon request.

### 4.2 US HIPAA Compliance (Health Insurance Portability and Accountability Act)

For deployments within the United States, LifeMetric acts as a Business Associate handling Protected Health Information (PHI).

1.  Business Associate Agreements (BAAs): LifeMetric must execute formal BAAs with:

Cloud hosting providers (e.g. AWS).

Target healthcare networks and Electronic Health Record (EHR) platforms (Epic, Cerner).

2.  Technical Safeguards (Security Rule § 164.312):

Access Control: Enforce unique user identification, RS256 JWT sessions, and automatic logout thresholds for clinician dashboards.

Transmission Security: Enforce TLS 1.3 as the minimum allowed transport layer protocol for all data in transit. Enforce AES-256 for all databases and S3 buckets at rest.

Audit Controls: Enable complete system audit trails. Utilize Postgres pgAudit to record all SQL queries touching vital sign tables and log all S3 object accesses via AWS CloudTrail.

---

## 5. Security & Data Lifecycle Architecture

Below is the data retention and encryption lifecyle designed to maintain strict privacy boundaries:

`

[ Raw Sensor Ingestion ] ──> Processed entirely in RAM ──> Dropped instantly (0ms retention)

│

▼

[ Behavior Metadata Event ] ──> Encrypted via TLS 1.3 ──> Transmitted to API Gateway

│

▼

[ Database / RDS Postgres ] ──> Encrypted via AES-256 ──> Retained under 90-day rolling cycle

│

▼

[ FHIR Observation Write ] ──> Auth via OAuth 2.0 ──> Written to EHR (Epic / Cerner)

`

1.  Local Memory Zeroing: On the physical device, raw mmWave registers and thermal grids are cleared from RAM immediately following local neural network inference execution.

2.  AWS Key Management Service (KMS): Database encryption keys are managed and rotated automatically every 90 days via AWS KMS KEK envelopes.

3.  Anonymized Telemetry Routing: Outbound telemetry packets do not contain patient names or medical histories; they are indexed strictly using a pseudonymous Device UUID. Patient identity mapping is resolved only within the hospital's on-premise EHR environment or secure FHIR API gateway.